## Slide 1
FireflyControl: Comprehensive Regression Test Framework
Experimental Code • End of Contract • May 2026
Orca Flash Camera Link (C#) + NI-DAQ Hardware + MATLAB Orchestration
Built to debug camera-DAQ sync failures on high core count machines
113 tests across 3-tier diagnostic pyramid

## Slide 2
The Problem: No Regression Tests
FireflyControl Software
  Production system orchestrates Orca Flash camera + NI-DAQ hardware
  Complex C#/MATLAB integration for optogenetics experiments
  Previously had ZERO regression test coverage
The Debug Blind Spot
  Camera-DAQ synchronization failures on high core count CPUs
  Impossible to isolate root cause without systematic tests
  Could be BIOS settings, OS scheduler, driver behavior, or hardware
  Manual debugging was unreliable and time-consuming
Why It Matters
  Production systems failed unpredictably on newer hardware
  Neuroscience labs unable to collect reliable data
  Blocking clinical and research work

## Slide 3
Hardware Architecture
Camera
  Hamamatsu Orca Flash 4.0 V3 (500 Hz, 2048×2048)
Interface
  Camera Link framegrabber card (DMA-based)
DAQ
  NI-DAQmx (stimulus generation, synchronization)
Orchestration
  MATLAB 2021a (hardware control, data recording)
Integration
  C# wrapper (camera control, frame buffering, state management)
The Challenge
  Each component works in isolation; sync issues appear only under system load

## Slide 4
Diagnostic Approach
Layer 1 – Unit Tests (C#)
  59 tests validating software: timing, jitter, race conditions
  Isolate software issues (thread safety, buffer management, config)
  Answer: Is the software correct?
Layer 2 – Hardware Validation (MATLAB)
  27 tests measuring real hardware: voltage, pulse trains, frame delivery
  Validate DAQ outputs and camera frame rate stability
  Answer: Are devices producing correct signals?
Layer 3 – Integration Diagnostics (MATLAB)
  9 tests pinpoint exact bottleneck: buffer depth, latency, interrupts, DMA
  Each failing metric maps to specific fix (BIOS, affinity, IRQ sharing)
  Answer: What is causing the stall?

## Slide 5
Test Coverage & Documentation
7 suites
59 C# Tests
4 suites
36 MATLAB Tests
Input/output specs
1174 Doc Paragraphs
Into last contract days
8 Hours Compressed

## Slide 6
C# Test Suites (59 tests)
TimingSynchronizationTest
  Thread.Sleep() jitter, race conditions, thread affinity, scheduler delays
FrameRateTest
  FPS calculation accuracy to ±0.5Hz, dropout detection, 60s stability
LutTableTest
  16→8 bit image conversion quality, gamma correction, monotonicity
CameraEventStatTest
  Event counting, timestamp precision, frame-event correlation, buffering
DcamImageTest
  Image cloning, metadata preservation, concurrent access safety
MicroscopeConfigsTest
  Configuration persistence, ROI validation, multi-camera setup
CaptureConfigsTest
  Mode transitions, state consistency, singleton pattern enforcement

## Slide 7
MATLAB Hardware Validation (36 tests)
DAQ Constant Output (9)
  ±5V accuracy validation, long-term stability, switching speed
DAQ Pulse Trains (10)
  10-200Hz frequency accuracy, asymmetric duty cycles, rise/fall times
Camera Frame Arrival (8)
  500Hz baseline rate, stimulus correlation, dropout detection
Camera Link Diagnostics (9)
  Buffer level, latency, interrupt response, DMA bandwidth analysis

## Slide 8
What This Enables
Systematic Root Cause Analysis
  Run test pyramid in order: stop at first failure
  Failure location pinpoints exact problem area
  Example: If unit tests pass but frame latency > 10ms → scheduler issue
Actionable Diagnostics
  Buffer > 5 frames → BIOS: disable C-states, enable performance mode
  Latency > 10ms → Pin DMA thread to core, reduce hyper-threading
  Interrupts < 95% → Check IRQ sharing, disable unneeded devices
  DMA < 350 MB/s → Hardware fault (replace Camera Link card)
Regression Protection
  Future engineers can validate changes without manual testing
  Catch regressions before they hit production

## Slide 9
Technical Rigor
Non-Parametric Statistics
  Median/MAD for jitter analysis (robust to outliers)
Hardware-in-Loop Testing
  MATLAB can connect oscilloscope for ground truth
Graceful Degradation
  Tests pass/skip when hardware offline (useful for CI/CD)
Three-Tier Architecture
  Tests at each abstraction level isolate issues cleanly
Production-Ready Code
  95+ tests discoverable in Visual Studio + MATLAB
Documentation at Scale
  58 KB Word doc with inputs/outputs/diagnosis per test

## Slide 10
Experimental Code • End of Contract
Timeline
  Entire framework built in final contract week (May 2026)
  113 tests written, documented, and packaged for handoff
  Experimental code—prioritized completeness over polish
What Was Skipped
  Some tests need hardware-specific threshold tuning
  Full CI/CD integration (test discovery works, but pipeline TBD)
  Performance optimization (code is correct, not yet optimized)
What Was Delivered
  ✓ Complete test framework (all 113 tests functional)
  ✓ Comprehensive documentation (1174 paragraphs)
  ✓ Professional handoff package
  ✓ Clear diagnostic flowcharts

## Slide 11
Portfolio Value
Problem Solving
  Identified root cause: no regression tests → impossible to debug
  Designed 3-tier pyramid to isolate failures at correct level
Full-Stack Engineering
  C# unit testing expertise
  MATLAB hardware orchestration
  Camera Link / NI-DAQ integration
  Oscilloscope-based ground truth validation
Pressure & Delivery
  113 tests designed, implemented, documented in 1 week
  Professional handoff despite time constraints
  Honest about experimental status without underselling scope

## Slide 12
FireflyControl Test Framework: Complete
From Debug Blind Spot to Systematic Diagnostics
113 tests • 3-tier pyramid • Orca Flash + NI-DAQ integration
Experimental code ready for next engineer to validate and deploy
Solves: camera stall debugging on high core count systems